## Recording equipment
Recording equipment
Type | ID | Product name | Model (REF/No/Version) | Vendor / Maker | Citation | Sampling frequency | Sampling
LVAD | R-01.1 | HeartWare Inc. FL 33014, USA | ? | HeartWare, Medtronic |  | ? | LVAD monitor
LVAD monitor | R-01.2 | HeartWare Monitor | REF 1521 1521DE HVAD Monitor | HeartWare, Medtronic | HeartWare Inc. FL 33014, USA | 1 datapoint per 15 min. | Log file exported to dedicated USB memory stick
LVAD controller | R-01.3 | HeartWare Controller | REF 1400 | HeartWare, Medtronic | HeartWare Inc. FL 33014, USA | ? | Analouge measurement of power, sampled at unknown rate. Flow rate derived from power measurement and user-set hematocrit value. Data tranferred to HeartWare monitor.
3-component accelerometer + gyroscope + magnetometer | R-02.1 | - | - | Høgskolen i Vestfold / Magnus Krogh, Cardiaccs | Cardiaccs Oslo, Norway | 700 Hz in sensor; 520-540 on laptop | Digital sensor signal to laptop with streaming software, or to DAC and then into PowerLab and LabChart
Digital-to-analouge converter (DAC) for 3-component accelerometer + gyroscope + magnetometer sensor signals | R-02.2 | - | - | Magnus Krogh, Cardiaccs | Cardiaccs Oslo, Norway | - | Up to 700 Hz input signal sampling rate; 2000 Hz sampling in LabChart
3-component accelerometer | R-03.1 | Cardisence | CS1 | Cardiaccs | Cardiaccs Oslo, Norway | 700Hz in sensor; 520-540 Hz in Teguar | Digital sensor signal to CardiBox and then to Teguar monitor and sampling software, or to DAC and then into PowerLab and LabChart
Signal isolation device for CS1 acceleromter and for ECG signal | R-03.2 | Cardibox | REF CB1-1 | Cardiaccs | Cardiaccs Oslo, Norway | 700Hz in sensor; 520-540 Hz in Teguar | Digital sensor signals from sensor are separated and isolated (demultiplexed), for which they og into Teguar monitor or to PC USB port and sampling software
Monitor for CS1 and ECG data | R-03.3 | Data aquistion unit (Medical Panel PC) | AU-1 (Model TM-4433-10) | Cardiaccs (Teguar) | Cardiaccs Oslo, Norway | 700Hz in sensor; 520-540 Hz in Teguar | Display and sampling to USB of CS1 and ECG data. Data is automatically stored to USB when memory stick is inserted in the USB port. (Only one sensor and USB output is currently supported.)
Digital-to-analouge converter (DAC) for CS1 sensor accelerometer signal | R-03.4 | CB1 Research DAC | CB1R | Cardiaccs | Cardiaccs Oslo, Norway | - | Up to 1500 Hz input signal content; 4000 Hz sampling in LabChart
Pressure tranducer | R-04 | Xtrans | BDC903-559NK REF 75.1216 | Codan |  | Setting in LabChart | Signal into Nihon Cohden, analouge signal in output channel and BP port into PowerLab + LabChart
Pressure tranducer | R-04 | Xtrans | BDC903-559NK REF 75.1216 | Codan |  | Setting in LabChart | Signal into Nihon Cohden, analouge signal in output channel and BP port into PowerLab + LabChart
Pressure tranducer for analouge signal | R-05 | - | - | Fred-Johan Pettersen, Med. Tek., OUS |  | Setting in LabChart | Analouge signal in output channel and BP port into PowerLab + LabChart
Driveline voltmeter/amperemeter | R-06 | - | - | Fred-Johan Pettersen, Med. Tek., OUS |  | Setting in LabChart | PowerLab + LabChart
PowerLab data acquisition (DAQ) device | R-07.1 | PowerLab 16/35 | Model PL3516 | AD Instruments | AD Instruments New South Wales, Australia | 400kS/s ADC giving a maximum per channel sampling rate of 200 kS/s | Data signal into analouge input channels, streamed digital data to LabChart and sampled at user-set sampling rate.
LabChart data acqusition software | R-07.2 | LabChart | Version ? | AD Instruments |  | Up to 200 kS/s | Streamed and stored data digitally sampled in PowerLab, at user-set sampling rate.
Ultrasound, clip-on, monitor | R-08.1 | System M | M4 | Spectrum Medical | Spectrum Medical Ltd. Gloucester, United Kingdom | 1 Hz | Signal from port 1 and port 2 (or port 3) are stored and exported to SD memory card at 1 Hz.
Ultrasound, clip-on, probe | R-08.2 | System M | H9XLM | Spectrum Medical | Spectrum Medical Ltd. Gloucester, United Kingdom | ? | Signal to port 1 and port 2
Matlab data processing software | R-09 | Matlab | 2019B | MathWorks | 1. MATLAB. 9.7.0.1190202 (R2019b). Natick, Massachusetts: The MathWorks Inc.; 2018. (see http://citebay.com/how-to-cite/matlab/) | - | Data resampled by Matlab's retime function
Ultrasound, echo probe | R-10 | ? | ? | ? |  | ? | ?
Millar pressure transducer catheter | R-11 | Mikro-Cath™ Diagnostic Pressure Catheter? | ? | Millar |  | 1000 Hz | Catheter to measure pressure. PowerLab + LabChart
Swan-Ganz catheter | R-12.1 | Swan-Ganz Thermodilution Catheter | ? | Edwards | Edwards Lifesciences Corporation Irvine, CA 92614, USA | ? | Pulmonary artery catheterization (PAC) or reight ventricle catheterization (RHC) to monitor hemodynamics. Data to Swan-Ganz monitor, exported to USB?
Swan-Ganz monitor | R-12.2 | ? | ? | Edwards |  | ? | Data to monitor, exported to USB?
Camera, 360 deg | R-13 | ? | ? | ? |  | x fps | Camera
Camera, action | R-14 | ? | ? | GoPro |  | x fps | Camera
X-ray machine | R-15 | ? | ? | ? |  | 30 fps (max) for balloon intervention, lower fps for catherter positioning guidence | Data stored in ?, exported by ???
Nihon Cohden bedside monitor | R-16 | Bedside monitor | Model BSM-1763 Ver 01-13 | Nihon Cohden | Nihon Cohden Corporation Tokyo, Japan | ? | Input signal from pressure transducer in multi port. Data streamed as analouge signal out in ECG/BP OUT port
Nihon Cohden monitor | R-17 |  |  | Nihon Cohden | Nihon Cohden Corporation Tokyo, Japan | ? | Input signal from pressure transducer in multi port. Data streamed as analouge signal out in ECG/BP OUT port

## Model equipment
Catheters for thrombosis simulation
Category | ID | Description | Introducer size | Max guide wire thickness | Vendor | Name | Model (REF)
PCI, 2mm x 28mm | C-1 | Coronary stent system, everolimus eluting, RX | 5Fr |  | Abbott | Xience Pro | 1076200-28
PCI, 2.5mm x 8mm | C-2 | Coronary stent system, everolimus eluting, RX | 5Fr |  | Abbott | Xience Pro | 1076225-08
PCI, 4mm x 20mm Papyrus | C-3 | Covered coronary stent system | 5Fr | 0.014" (0.36mm) | BIOTRONIK | PK Papyrus | 369389
PCI, 4mm x 20mm | C-4 | PTCA balloon dilatation, rapid exchange | 6Fr | 0.014" (0.36mm) | Cordis / Kaneka Corporation | Raiden 3 | RD3-20-400
PCI, 4.5mm x 12mm | C-5 | Coronary dilatation, RX |  | 0.014" (0.36mm) | Abbott | TREK | 1012279-12
PCI, 4.5mm x 15mm | C-6 | PTCA balloon dilatation, non-compliant | 6Fr | 0.014" (0.36mm) | Cordis / Kaneka Corporation | NEON NC | BN-45015
PCI, 4.5mm x 20mm | C-7 | PTCA balloon dilatation, non-compliant | 6Fr | 0.014" (0.36mm) | Cordis / Kaneka Corporation | NEON NC | BN-45020
PCI, 6mm x 8mm | C-8 | PTCA balloon dilatation, non-compliant | 6Fr | 0.014" (0.36mm) | Boston Scientific | NC Emerge, MONORAIL | H7493927608600
PCI, 6mm x 18mm | C-9 | Renal Stent System, RX | 6Fr | ? | Abbott | HLINK (Herkulink?) ELITE | Order no 1011498
PCI, 6mm x 20mm | C-10 | PTA | 5Fr | 0.035" (0.89mm) | Abbott | Armada 35 | B1080-020
PCI, 6mm x 39mm | C-11 | PTCA, paclitaxel-eluting (drug-eluting) | 4Fr | 0.014" (0.36mm) | Aachen Resonance | ELUTAX | ELUTAX39600
PCI, 8mm x 20mm | C-12 | PTA | ? | ? | Abbott | Armada 35 | B-2080-020
PCI, 8mm x 30mm | C-13 | PTA | 6Fr | 0.035" (0.89mm) | Abbott | Fox | A13008
PCI, 11mm x 30mm | C-14 | Percutaneous tranluminal valvuloplasty | 5Fr | 0.025" | NuMed | TYSHAK II | SO017
RHC, 11mm | C-15 | Pumonary wedge catheter | 7Fr |  | Arrow | Balloon Wedge Pressure Catheter | AI-07127
Note on C13: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: https://www.cardiovascular.abbott/us/en/hcp/products/peripheral-intervention/rx-herculink-elite-renal-stent-system.html
Note on C18: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: http://www.bisusa.org/sites/bbrauninterventionalsystems.com/files/documents/ifu105barev01tyshak-ii-ptv.pdf